Cold Chain Checklist for Immunisation Providers

The following checklist has been developed to help Immunisation Providers assess cold chain compliance. Please call or email the Northern Sydney Public Health Unit on 9485 6975 nslhd-phuimmunisation@health.nsw.gov.au if you require any assistance.

Refrigerator

  - Is a vaccine specific model (the use of domestic fridges is no longer supported)

  - Is serviced annually and service report is kept on file

  - Is plugged in directly to a wall socket and clearly labelled not to disconnect or turn off

  - Front feet are adjusted higher than the back to allow a slight tilt to ensure door closes freely

  - There is an appropriate gap between the vaccines and the walls of the refrigerator

  - Can safely store the required volume of vaccines, including times of increased demand

  - Has no food or drink kept inside – is for temperature sensitive medications only

  - Vaccine stocks are kept to a minimum for business requirements - do not stockpile.

  - Vaccine stocks are rotated, putting older stock to the front and newer stock to the back

  - Vaccine stocks are kept in their original packaging, protected from light

  - Solid door refrigerator has a map/guide on the door to where vaccines are located inside

  - Alternate monitored fridge storage is available in case of vaccine fridge malfunction

  - Current, minimum and maximum temperatures are checked and recorded twice daily.

  - Fridge temperatures are reset after each check and/or after restocking

  - All deviations outside +2°C to +8°C (excluding ≤ 12°C for ≤ 15 minutes) reported to PHU ASAP

  - A copy of the Cold Chain Breach Protocol is displayed on the fridge https://www.health.nsw.gov.au/immunisation/Pages/ccb-protocol.aspx

Coolers and ice/gel packs

- Coolers are solid-walled insulated containers with a tightly fitting lid (no smaller than 10L)

- There are sufficient numbers of coolers to contain all stocks of vaccines in case of a CCB

- Sufficient number of ice/gel packs available to keep cooler temperatures in range during CCB

- There is enough insulating material for each cooler e.g. bubble wrap

Minimum/maximum thermometers

- Battery operated min/max are in use as a backup for the fridge in case of power outage

- There are sufficient numbers for each fridge and each cooler

- Temperature probes are placed between vaccine stock, ideally in an old vaccine box

- Battery has been changed annually and date of change recorded on the device

- Slush test (refer to Strive for 5 for instructions) has been attended annually

Data logger

- An inbuilt/portable data logger is in use for all fridges storing vaccines

- The data logger is placed amongst vaccine stock to best replicate vaccine temperature

- Set data logger to record continuously at 5 minutely intervals, even during power failure

- The data logger and/or battery is changed as per manufacturer’s instructions

- Data logger report is downloaded, saved and reviewed weekly

Staff Education

- All staff are aware of the importance of cold chain management for vaccines

- A staff member and back up staff has been assigned for vaccine cold chain management

- All staff involved in the cold chain have completed the NSW Health Vaccine storage and Cold chain management online training module annually https://nswhealth.seertechsolutions.com.au/public_content/HETICP/HETI/CCMWebv3/story_html5.html

- A vaccine storage self-audit (refer to Strive for 5 template) has been completed annually